## Hazards
No. | HML-Style Hazard Description
H1 | Button pressed(Being pressed:Button)(Press button<Wrong moment>:Patient)
Natural Language Hazard Description
Press the button in a non-scheduled time
No. | Pre-Initiating event for H1 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Accidentaly | 0.25 | 0.25 | 0.0625 | Accept | M 1
Cause 2 | Misinterpretation of instructions | 0.25 | 0.75 | 0.1875 | Denied | M 2
No. | HML-Style Hazard Description
H2 | Button pressed(Being pressed:Button)(Press button<Not able to press>:Patient)
Natural Language Hazard Description
Patient unable to press the button
No. | Pre-Initiating event for H2 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Unfamiliarity with system operation | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H3 | Button pressed(Being pressed:Button)(Press button<Not pressed on time>:Patient)
Natural Language Hazard Description
Patient do not press the button in the scheduled time
No. | Pre-Initiating event for H3 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Patient forgets | 0.5 | 0.5 | 0.25 | Denied | M 1
Cause 2 | Misinterpretation of doctor prescription | 0.25 | 0.75 | 0.1875 | Denied | M 2
No. | HML-Style Hazard Description
H4 | Button pressed(Being pressed:Button)(Being pressed<Physical damage>:Button)
Natural Language Hazard Description
Damage of the button system
No. | Pre-Initiating event for H4 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Deterioration due to long term use | 0.25 | 0.5 | 0.125 | Accept | M 1
Cause 2 | Deterioration due to accidental event | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H5 | Select dose(Being modified:Potentiometer)(Vary dosage<Wrong dosage>:Doctor)
Natural Language Hazard Description
Doctor prescribes the wrong dose
No. | Pre-Initiating event for H5 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Error in the programming of the device | 0.25 | 0.75 | 0.1875 | Denied | M 1
Cause 2 | Doctor's calculation error | 0.25 | 0.75 | 0.1875 | Denied | M 2
No. | HML-Style Hazard Description
H6 | Select dose(Being modified:Potentiometer)(Being modified<Physical damage>:Potentiometer)
Natural Language Hazard Description
Damage of the potentiometer system
No. | Pre-Initiating event for H6 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Deterioration | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H7 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Potentiometer)
Natural Language Hazard Description
Sending wrong information
No. | Pre-Initiating event for H7 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | System deterioration | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H8 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Encoder)
Natural Language Hazard Description
Sending wrong information
No. | Pre-Initiating event for H8 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | System deterioration | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H9 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Button)
Natural Language Hazard Description
Sending wrong information
No. | Pre-Initiating event for H9 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | System deterioration | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H10 | Information transfer(Receive information:Processor)(Receive information<Unable to receive>:Processor)
Natural Language Hazard Description
Processor unable to receive information
No. | Pre-Initiating event for H10 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Failure in the transmission system (wires, electronic components...) | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H11 | Information transfer(Receive information:Processor)(Receive information<Misunderstanding information>:Processor)
Natural Language Hazard Description
Processor misunderstanding good information
No. | Pre-Initiating event for H11 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Failure in the communication protocol | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H12 | Electricity consumption(Electricity consumer:Processor)(Electricity source<Wrong power supply>:Power supply)
Natural Language Hazard Description
Error in the power supply
No. | Pre-Initiating event for H12 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Too much / too little voltage | 0.25 | 0.5 | 0.125 | Accept | M 1
Cause 2 | Too much / too little current | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H13 | Control code(Code consumer:Processor)(Code source<CodeErrors>:Firmware)
Natural Language Hazard Description
Code errors leading to undesired behavior
No. | Pre-Initiating event for H13 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Any kind of coding errors | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H14 | Motor sensing(Being sensed:Motor)(Sense position<Alignment>:Encoder)
Natural Language Hazard Description
Missalignment of encoder sensor and motor leading to wrong signals
No. | Pre-Initiating event for H14 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Wrong manufacturing | 0.25 | 0.5 | 0.125 | Accept | M 1
Cause 2 | Misalignment due to rapid acceleration or hit | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H15 | Motor sensing(Being sensed:Motor)(Sense position<Deterioration>:Encoder)
Natural Language Hazard Description
Physical damage on sensor
No. | Pre-Initiating event for H15 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Deterioration due to long term use | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H16 | Drug delivery(Being pushed:Drug)(Push drug<Rotating too much>:Motor)
Natural Language Hazard Description
Too much rotation of the motor
No. | Pre-Initiating event for H16 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Motor rotation calibration failure | 0.25 | 0.75 | 0.1875 | Denied | M 1
Cause 2 | Wrong control signal received | 0.25 | 0.75 | 0.1875 | Denied | M 2
No. | HML-Style Hazard Description
H17 | Drug delivery(Being pushed:Drug)(Being emptied<Insuffiecient drug>:Cartridge)
Natural Language Hazard Description
No drug or insufficient drug in the cartridge
No. | Pre-Initiating event for H17 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Patient did not refill the cartridge | 0.25 | 0.5 | 0.125 | Accept | M 1
No. | HML-Style Hazard Description
H18 | Drug delivery(Being pushed:Drug)(Being emptied<Interaction with drug>:Cartridge)
Natural Language Hazard Description
chemical/physical interaction with drug in the cartridge
No. | Pre-Initiating event for H18 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Wrong combination of device material and drug | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H19 | Drug delivery(Being pushed:Drug)(Being emptied<Obstruction>:Cartridge)
Natural Language Hazard Description
Obstruction preventing drug release
No. | Pre-Initiating event for H19 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Dust or external elements in the cartridge exit | 0.5 | 0.5 | 0.25 | Denied | M 1
No. | HML-Style Hazard Description
H20 | Drug delivery(Being pushed:Drug)(Being emptied<Incorrect position>:Cartridge)
Natural Language Hazard Description
Incorrect position of cartridge
No. | Pre-Initiating event for H20 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Incorrect placement of cartridge by the patient | 0.5 | 0.25 | 0.125 | Accept | M 1
Cause 2 | Incorrect placement due to mechanism deterioration | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H21 | Drug delivery(Being pushed:Drug)(Being pushed<Incorrect drug>:Drug)
Natural Language Hazard Description
Incorrect drug
No. | Pre-Initiating event for H21 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Patient introduces incorrect drug to the device | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H22 | Drug delivery(Being pushed:Drug)(Being pushed<Physical properties>:Drug)
Natural Language Hazard Description
Non-suitable physical properties of drug for the system
No. | Pre-Initiating event for H22 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Drug not suitable for the mechanism of the device | 0.25 | 0.75 | 0.1875 | Denied | M 1
No. | HML-Style Hazard Description
H23 | Drug delivery(Being pushed:Drug)(Being filled<Obstruction2>:Dispenser)
Natural Language Hazard Description
Obstruction preventing supply
No. | Pre-Initiating event for H23 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Dust or external elements obstructing the entrance | 0.5 | 0.5 | 0.25 | Denied | M 1
No. | HML-Style Hazard Description
H24 | Drug delivery(Being pushed:Drug)(Being filled<Incorrect position 2>:Dispenser)
Natural Language Hazard Description
Incorrect position of dispenser
No. | Pre-Initiating event for H24 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Incorrect placement by the patient | 0.5 | 0.25 | 0.125 | Accept | M 1
Cause 2 | Incorrect placement due to deterioration of mechanism | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H25 | Drug intake(Being consumed:Drug)(Consume drug<Intake problem>:Patient)
Natural Language Hazard Description
Patient not able to consume drug
No. | Pre-Initiating event for H25 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Swallowing problems | 0.25 | 0.75 | 0.1875 | Denied | M 1
Cause 2 | Drug fatality before intake | 0.25 | 0.5 | 0.125 | Accept | M 2
No. | HML-Style Hazard Description
H26 | Electricity consumption(Electricity consumer:Motor)(Electricity source<Wrong power supply>:Power supply)
Natural Language Hazard Description
Error in the power supply
No. | Pre-Initiating event for H26 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Already covered | 0.25 | 0.25 | 0.0625 | Accept | M 1
No. | HML-Style Hazard Description
H27 | Electricity consumption(Electricity consumer:Encoder)(Electricity source<Wrong power supply>:Power supply)
Natural Language Hazard Description
Error in the power supply
No. | Pre-Initiating event for H27 | Probability | Severity | RPN | Risk | Rq. | Mitigation
Cause 1 | Already covered | 0.25 | 0.25 | 0.0625 | Accept | M 1

## Causes
No. | HML-Style Hazard Description | Natural Language Hazard Description | Cause | Probability | Severity | RPN | Risk | Rq. | Mitigation
cause 1, H:1 | Button pressed(Being pressed:Button)(Press button<Wrong moment>:Patient) | Press the button in a non-scheduled time | Accidentaly | 0.25 | 0.25 | 0.0625 | Accept | M:1
cause 1, H:26 | Electricity consumption(Electricity consumer:Motor)(Electricity source<Wrong power supply>:Power supply) | Error in the power supply | Already covered | 0.25 | 0.25 | 0.0625 | Accept | M:1
cause 1, H:27 | Electricity consumption(Electricity consumer:Encoder)(Electricity source<Wrong power supply>:Power supply) | Error in the power supply | Already covered | 0.25 | 0.25 | 0.0625 | Accept | M:1
cause 1, H:2 | Button pressed(Being pressed:Button)(Press button<Not able to press>:Patient) | Patient unable to press the button | Unfamiliarity with system operation | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:4 | Button pressed(Being pressed:Button)(Being pressed<Physical damage>:Button) | Damage of the button system | Deterioration due to long term use | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 2, H:4 | Button pressed(Being pressed:Button)(Being pressed<Physical damage>:Button) | Damage of the button system | Deterioration due to accidental event | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 1, H:9 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Button) | Sending wrong information | System deterioration | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:10 | Information transfer(Receive information:Processor)(Receive information<Unable to receive>:Processor) | Processor unable to receive information | Failure in the transmission system (wires, electronic components...) | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:11 | Information transfer(Receive information:Processor)(Receive information<Misunderstanding information>:Processor) | Processor misunderstanding good information | Failure in the communication protocol | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:12 | Electricity consumption(Electricity consumer:Processor)(Electricity source<Wrong power supply>:Power supply) | Error in the power supply | Too much / too little voltage | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 2, H:12 | Electricity consumption(Electricity consumer:Processor)(Electricity source<Wrong power supply>:Power supply) | Error in the power supply | Too much / too little current | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 1, H:14 | Motor sensing(Being sensed:Motor)(Sense position<Alignment>:Encoder) | Missalignment of encoder sensor and motor leading to wrong signals | Wrong manufacturing | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 2, H:14 | Motor sensing(Being sensed:Motor)(Sense position<Alignment>:Encoder) | Missalignment of encoder sensor and motor leading to wrong signals | Misalignment due to rapid acceleration or hit | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 1, H:15 | Motor sensing(Being sensed:Motor)(Sense position<Deterioration>:Encoder) | Physical damage on sensor | Deterioration due to long term use | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:17 | Drug delivery(Being pushed:Drug)(Being emptied<Insuffiecient drug>:Cartridge) | No drug or insufficient drug in the cartridge | Patient did not refill the cartridge | 0.25 | 0.5 | 0.125 | Accept | M:1
cause 1, H:20 | Drug delivery(Being pushed:Drug)(Being emptied<Incorrect position>:Cartridge) | Incorrect position of cartridge | Incorrect placement of cartridge by the patient | 0.5 | 0.25 | 0.125 | Accept | M:1
cause 2, H:20 | Drug delivery(Being pushed:Drug)(Being emptied<Incorrect position>:Cartridge) | Incorrect position of cartridge | Incorrect placement due to mechanism deterioration | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 1, H:24 | Drug delivery(Being pushed:Drug)(Being filled<Incorrect position 2>:Dispenser) | Incorrect position of dispenser | Incorrect placement by the patient | 0.5 | 0.25 | 0.125 | Accept | M:1
cause 2, H:24 | Drug delivery(Being pushed:Drug)(Being filled<Incorrect position 2>:Dispenser) | Incorrect position of dispenser | Incorrect placement due to deterioration of mechanism | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 2, H:25 | Drug intake(Being consumed:Drug)(Consume drug<Intake problem>:Patient) | Patient not able to consume drug | Drug fatality before intake | 0.25 | 0.5 | 0.125 | Accept | M:2
cause 2, H:1 | Button pressed(Being pressed:Button)(Press button<Wrong moment>:Patient) | Press the button in a non-scheduled time | Misinterpretation of instructions | 0.25 | 0.75 | 0.1875 | Denied | M:2
cause 2, H:3 | Button pressed(Being pressed:Button)(Press button<Not pressed on time>:Patient) | Patient do not press the button in the scheduled time | Misinterpretation of doctor prescription | 0.25 | 0.75 | 0.1875 | Denied | M:2
cause 1, H:5 | Select dose(Being modified:Potentiometer)(Vary dosage<Wrong dosage>:Doctor) | Doctor prescribes the wrong dose | Error in the programming of the device | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 2, H:5 | Select dose(Being modified:Potentiometer)(Vary dosage<Wrong dosage>:Doctor) | Doctor prescribes the wrong dose | Doctor's calculation error | 0.25 | 0.75 | 0.1875 | Denied | M:2
cause 1, H:6 | Select dose(Being modified:Potentiometer)(Being modified<Physical damage>:Potentiometer) | Damage of the potentiometer system | Deterioration | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:7 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Potentiometer) | Sending wrong information | System deterioration | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:8 | Information transfer(Receive information:Processor)(Send information<Wrong information>:Encoder) | Sending wrong information | System deterioration | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:13 | Control code(Code consumer:Processor)(Code source<CodeErrors>:Firmware) | Code errors leading to undesired behavior | Any kind of coding errors | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:16 | Drug delivery(Being pushed:Drug)(Push drug<Rotating too much>:Motor) | Too much rotation of the motor | Motor rotation calibration failure | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 2, H:16 | Drug delivery(Being pushed:Drug)(Push drug<Rotating too much>:Motor) | Too much rotation of the motor | Wrong control signal received | 0.25 | 0.75 | 0.1875 | Denied | M:2
cause 1, H:18 | Drug delivery(Being pushed:Drug)(Being emptied<Interaction with drug>:Cartridge) | chemical/physical interaction with drug in the cartridge | Wrong combination of device material and drug | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:21 | Drug delivery(Being pushed:Drug)(Being pushed<Incorrect drug>:Drug) | Incorrect drug | Patient introduces incorrect drug to the device | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:22 | Drug delivery(Being pushed:Drug)(Being pushed<Physical properties>:Drug) | Non-suitable physical properties of drug for the system | Drug not suitable for the mechanism of the device | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:25 | Drug intake(Being consumed:Drug)(Consume drug<Intake problem>:Patient) | Patient not able to consume drug | Swallowing problems | 0.25 | 0.75 | 0.1875 | Denied | M:1
cause 1, H:3 | Button pressed(Being pressed:Button)(Press button<Not pressed on time>:Patient) | Patient do not press the button in the scheduled time | Patient forgets | 0.5 | 0.5 | 0.25 | Denied | M:1
cause 1, H:19 | Drug delivery(Being pushed:Drug)(Being emptied<Obstruction>:Cartridge) | Obstruction preventing drug release | Dust or external elements in the cartridge exit | 0.5 | 0.5 | 0.25 | Denied | M:1
cause 1, H:23 | Drug delivery(Being pushed:Drug)(Being filled<Obstruction2>:Dispenser) | Obstruction preventing supply | Dust or external elements obstructing the entrance | 0.5 | 0.5 | 0.25 | Denied | M:1